NAME OF THE EQUIPMENT: Chemiluminiscence Immunoassay Analyser

TECHNICAL REQUIREMENT:

Intended use:

Fully automated clinical Chemiluminiscence Immunoassay Analyser system for diagnostic patient samples

Can quote for more than one model: Quote only for reagent rental basis (not for outright purchase)

No. of Equipment required: 1

Equipment: For Composite and Microbiology lab

Cost / Test including inclusive cost of all reagent and consumables except controls and calibrators for all tests mentioned should be quoted in the financial bid

Technically accepted L1 will be identified based on the total Cost / test inclusive of consumable cost for 1 month (30 days).

Name of the Manufacturer:

Name of the Vendor:

| Sr. No. | SPECIFICATIONS | VENDOR’S OFFER (all rows to be filled in detail) |
| --- | --- | --- |
| 1 | Name of the Model: |  |
| 2 | Model No. of the Equipment (If the equipment is a combination of more than one equipment then list down the model no. pertaining to each) : |  |
| 3 | Year of Introduction Internationally |  |
| 4 | Year of Introduction in India |  |
| 5 | Specification: |  |
| 5.1 | System should be fully automated continuous loading analyzer with random access. |  |
| 5.2 | System should use Chemiluminescence technology with high sensitivity and linearity |  |
| 5.3 | System should have on board sample capacity of at least 60 or more samples at one time with provision of continuous loading of samples |  |
| 5.4 | System should have test throughput of 400 tests or more per hour from a single analysis unit |  |
| 5.5 | ## System should have provision of STAT samples. |  |
| 5.6 | ## System should have on board refrigeration facility for reagents with 40 positions or more, system should have capacity to control temperature and humidity while processing assays. |  |
| 5.7 | ## Assay reagents should be ready to use |  |
| 5.8 | ## System should have sample bar code reading facility |  |
| 5.9 | System software should have patient history recall option. |  |
| 5.10 | The system should be interfaced with the Hospital diagnostic information system for patient samples. |  |
| 5.11 | Interfacing with Hospital information system will be the vendor’s responsibility. |  |
| 5.12 | The system should have disposable tips or washable probes to avoid carryover |  |
| 5.13 | Footprint should not be greater than 7ft in width/length X 5 ft in depth (inclusive of monitor) |  |
| 5.14 | Expected water consumption per hour for full load not more than 10lit/hour preferred |  |
| 5.15 | To provide backup analyser as and when needed |  |
| 6 | Anticipated workload |  |
| 6.1 | Equipment 1: (Present sample load) |  |
|  | System should be able to process following parameters: (Mandatory) - Serum T3 (100samples/month) - Serum T4 (100samples/month) - Serum TSH (115samples/month) - Free T3 (25 samples/month) - Free T4 (25 samples/ month) - Serum Folate (110samples/month) - Serum B12 (120 samples/month) - Vitamin D (100samples/month) - Serum PSA (25samples/month) - Serum HCG (20samples/month) - Serum Ferritin (75samples/month) - Serum Estradiol (10 samples/ month) - C-peptide (15 samples/ month) - Insulin (15 samples/ month) - HIV Ag/Ab Combo – (200 samples/ month) - Anti HCV – (200 samples/ month) - HBs Ag Qualitative – (200 samples/ month) - Anti HBs – (50 samples/ month) - Anti HBc IgM – (50 samples/ month) - Anti HBc Total –( 50 samples/ month) |  |
| 6.2 | To also quote all other relevant parameters which may have smaller volume and will order as and when required (Optional) |  |
| 7 | Cost / Test |  |
| 7.1 | To provide Cost/ test inclusive of cost of all reagents and consumables except controls and calibrators (CPTI) and the methodology adopted for its calculation in the financial bid for tests indicated in 6.1(refer table) |  |
| 7.2 | L1 will be identified based on total of total cost / month of all tests with the indicated volumes for parameters listed as mandatory in 6.1 |  |
| 8 | Equipment which will be placed should be brand new. It should not be used or refurbished. |  |
| 9 | Suitable rating UPS should be provided with the system |  |
| 10 | US FDA and CE approved system |  |
| 10 | Any technological updates and upgrades should be installed as and when it is available. |  |
| 11 | Standard accessories (Which comes along with the machine, pricing of which is part of the equipment) |  |
| 12 | Essential accessories (Essential to carry out certain functions but comes at an extra cost. Prices to be quoted separately) |  |
| 13 | Optional accessories (List out all the optional accessories. Prices to be quoted separately in the Financial bid |  |
| 14 | Any Additional / salient features which you want to highlight pertaining to the model quoted |  |
| 15 | Essential consumables: |  |
| 15.1 | Provide List of consumables with their prices in the Financial bid. |  |
| 16 | Pre Installation Requirement |  |
| 17 | Installation: Company trained personnel as per end user’s satisfaction should do installation and impart training to end user and provide IQ, OQ, PQ of the equipment. |  |
| 18 | Training of staff: Provide necessary training to Institute’s staff / engineers in routine maintenance and fault finding. |  |
| 19 | Service support: |  |
| 19.1 | Complete operational manuals and technical information should be provided. Company should provide technical support for hardware and software, spares and consumables. |  |
| 19.2 | Repair and Maintenance The equipment should be entirely maintained by the company with periodic preventive maintenance, calibrations and break down repairs including spare parts. The company has to ensure uptime guarantee of 95% taking into consideration 313 working days in a year. |  |
| 19.3 | Response to break down at site: within 4 hrs |  |
| 19.4 | Provision for standby unit if downtime exceeds 48 hrs. |  |
| 20 | Regulatory Approvals if any. Details with copies of approvals |  |
| 21 | Safety requirements: It should follow International safety requirement. Please specify certification with certifying agency and country with copies of certificate. |  |
| 22 | Input power requirements: 230 volts +/- 10% , 50 Hz single phase or 400V/ 3 phase supply with inbuilt / optional protection for high / low voltage ( Please mention the actual power requirement) |  |
| 23 | User’s list: A list of installations with the address and contact numbers to be provided. (User list should be for the quoted model) |  |
| 25 | Multiple Models: If more then one model is offered, then they are to be separately filled and submitted according to the given format. |  |

*For Equipment on Placement basis, approximate monthly utilization, Cost per test and how the calculation is done to arrive at cost per test is to be indicated. Maintenance will be vendor’s responsibility for all placement equipment

Table:

| Sr. no. | Parameter (A) | Reagent cost/Test (B) | Cost of all Consumables / Test ( C ) | Cost/ Test including Reagent and Consumable Cost (B+C) |
| --- | --- | --- | --- | --- |